УДК 631.432:681.2

СИСТЕМАТИЧЕСКАЯ ПОГРЕШНОСТЬ ЗАВОДСКОЙ КАЛИБРОВКИ ДАТЧИКОВ ВЛАЖНОСТИ CS616 В НЕЗАСОЛЁННЫХ ПОЧВАХ РАЗНОГО ГРАНУЛОМЕТРИЧЕСКОГО СОСТАВА: ЭКСПЕРИМЕНТАЛЬНАЯ ОЦЕНКА И СТАТИСТИЧЕСКИЙ АНАЛИЗ

Вербицкий А. Ю., Приходько И. А., Болдырева Л. М., Кесафоти Х. Е.

Кубанский государственный аграрный университет имени И. Т. Трубилина, Краснодар, Россия

Аннотация

Обоснование. Рефлектометрические датчики CS616 (Campbell Scientific, США) широко применяются для непрерывного контроля объёмной влажности почвы (θ) при управлении орошением. Паспортный предел погрешности стандартной (заводской) калибровки датчика (±0,025 м³/м³) установлен преимущественно по измерениям в почвах класса sandy loam (по классификации USDA) и более лёгких, поэтому для тяжелосуглинистых и глинистых почв его соблюдение не гарантировано. В таких почвах электропроводность порового раствора и поверхностная электропроводность глинистых минералов увеличивают период выходного сигнала, что может приводить к систематическому завышению θ, нарастающему по мере приближения к насыщению. Для чернозёмов и лугово-чернозёмных почв Краснодарского края количественные оценки этой погрешности и влияния на неё плотности сложения и набухания в доступной литературе отсутствуют.

Цель. Экспериментально оценить систематическую погрешность заводской калибровки датчиков CS616 в незасолённых почвах трёх гранулометрических групп при двух уровнях плотности сложения и определить вклад гранулометрического состава, плотности сложения и их взаимодействия в изменчивость этой погрешности.

Материалы и методы. Двухфакторный эксперимент 3 × 2 (часть полного факторного эксперимента 3 × 2²); пипеточный метод Качинского с пирофосфатной подготовкой; пикнометрическое определение плотности твёрдой фазы; определение максимальной гигроскопической влажности; кондуктометрия водной вытяжки (1 : 5) и вытяжки из водонасыщенной пасты; лабораторная проверка заводской калибровки в колонках с почвой нарушенного сложения; термостатно-весовой метод (референтный); оценка средней систематической (MB) и среднеквадратической (RMSE) погрешностей; одновыборочный t-критерий; двухфакторный дисперсионный анализ; двухфакторный дисперсионный анализ без повторностей (уровень влажности — блок) и критерий Фридмана; линейный регрессионный анализ.

Результаты. В исследованных почвах (содержание физической глины, т. е. частиц < 0,01 мм: группа A — 10–20 %, B — 40–50 %, C — 50–65 %; удельная электропроводность вытяжки из водонасыщенной пасты ECe = 0,28–0,61 дСм/м; плотность сложения 1,35 и 1,52 г/см³) RMSE заводской калибровки (по средним значениям для трёх колонок) составила 0,006–0,014 м³/м³, что меньше принятого критерия 0,02 м³/м³. Наибольшее значение погрешности среди десяти анализируемых уровней влажности (+0,027 м³/м³, среднее по трём колонкам; группа C при плотности сложения 1,35 г/см³ и θ = 0,40 м³/м³) превысило паспортный предел погрешности датчика (±0,025 м³/м³) на 0,002 м³/м³, т. е. на величину, меньшую расширенной неопределённости референтного метода (0,005 м³/м³). При двух более влажных состояниях этого варианта, зарегистрированных, но не включённых в анализ (θ = 0,44 и 0,48 м³/м³), Δθ составила +0,031 и +0,038 м³/м³. RMSE возрастала с увеличением содержания физической глины: в среднем 0,0065; 0,0100 и 0,0130 м³/м³ в группах A, B и C соответственно. Значения MB во всех вариантах были положительными; в группе A MB статистически значимо не отличалась от нуля (+0,002…+0,004 м³/м³; p = 0,066 и 0,32), а в группах B и C датчик завышал влажность в среднем на 0,006–0,010 м³/м³ (одновыборочный t-критерий по десяти уровням влажности, p = 0,014–0,035, оценка описательная; систематическая составляющая неопределённости референтного метода (U = 0,005 м³/м³) не учтена; различие групп A и B при α = 0,05 чувствительно к округлению MB и RMSE). Значения MB по вариантам возрастали с увеличением максимальной гигроскопической влажности (линейная аппроксимация по шести вариантам: 0,00066 м³/м³ на 1 %; R² = 0,88; p = 0,006; по средним значениям трёх групп p = 0,07), но не менее тесно были связаны с ECe (R² = 0,96), поэтому вклады электропроводности порового раствора и поверхностной электропроводности разделить не удалось. При плотности сложения 1,35 г/см³ в интервале θ = 0,20–0,40 м³/м³ погрешность увеличилась на 0,008; 0,014 и 0,021 м³/м³ в группах A, B и C соответственно. В двухфакторном дисперсионном анализе (остаточная изменчивость включает зависимость погрешности от θ) различия между вариантами объясняли около 10 % общей изменчивости, а влияние гранулометрического состава (9 %) было статистически незначимым (F(2; 54) = 2,74; p = 0,074). Однако при сравнении групп на шести одинаковых уровнях θ (плотность сложения 1,35 г/см³) погрешность на каждом уровне возрастала в ряду A < B ≤ C (влияние гранулометрического состава в двухфакторном дисперсионном анализе без повторностей: F(2; 10) = 6,92; p = 0,013; оценка описательная). Влияние плотности сложения было малым и статистически незначимым (разность MB −0,001…−0,002 м³/м³; p = 0,43), тогда как модель комплексного показателя преломления (CRIM) предсказывает увеличение показаний примерно на 0,009–0,010 м³/м³; взаимодействие факторов не выявлено (p = 0,98). Почва группы C при плотности сложения 1,52 г/см³ набухала (объёмная деформация 8,3 %; максимальная влажность θmax = 0,460 м³/м³ превысила исходную пористость 0,415 м³/м³); при расчёте референтной влажности по фактическому объёму колонки дополнительная погрешность, связанная с набуханием, не выявлена (в этом сравнении влияние набухания нельзя отделить от влияния плотности сложения и ECe). Заводская калибровка CS616 пригодна для исследованных незасолённых почв (физическая глина 10–20 и 40–65 %), если допустима RMSE 0,02 м³/м³, однако в исследованных почвах с содержанием физической глины 40–65 % датчик с заводской калибровкой завышал влажность в среднем на 0,006–0,010 м³/м³ (до 0,027 м³/м³ при θ = 0,40 м³/м³), что целесообразно компенсировать калибровкой датчика для конкретной почвы или контролировать периодическими определениями влажности термостатно-весовым методом.

Ключевые слова: датчик CS616, объёмная влажность почвы, заводская калибровка, систематическая погрешность, гранулометрический состав, плотность сложения, электропроводность почвы, набухание, дисперсионный анализ, чернозём

SYSTEMATIC ERROR OF THE FACTORY CALIBRATION OF CS616 SOIL WATER CONTENT SENSORS IN NON-SALINE SOILS OF DIFFERENT TEXTURE: EXPERIMENTAL EVALUATION AND STATISTICAL ANALYSIS

Verbitsky A. Yu., Prikhodko I. A., Boldyreva L. M., Kesafoti Kh. E.

Kuban State Agrarian University named after I. T. Trubilin, Krasnodar, Russia

Abstract

Background. CS616 water content reflectometers (Campbell Scientific, USA) are widely used for continuous monitoring of volumetric soil water content (θ) in irrigation management. The manufacturer's accuracy specification for the standard (factory) calibration (±0.025 m³/m³) is based mainly on measurements in sandy loam (USDA classification) and coarser soils, so compliance with it is not guaranteed in heavy loam and clay soils. In such soils, the conductivity of the pore solution and the surface conductivity of clay minerals lengthen the output period of the sensor, which may cause a systematic overestimation of θ that increases as the soil approaches saturation. For the chernozems and meadow-chernozem soils of the Krasnodar Krai, no quantitative estimates of this error or of the effects of bulk density and swelling on it are available in the literature.

Purpose. To evaluate experimentally the systematic error of the CS616 factory calibration in non-saline soils of three texture groups at two bulk density levels and to quantify the contributions of texture, bulk density and their interaction to its variability.

Materials and methods. Two-factor 3 × 2 experiment (part of a full 3 × 2² factorial design); Kachinskii pipette method with pyrophosphate pretreatment; pycnometric determination of particle density; determination of maximum hygroscopic water content; electrical conductivity measurement of 1:5 soil–water and saturated-paste extracts; laboratory evaluation of the factory calibration in repacked soil columns; oven-drying gravimetric reference method; mean bias (MB) and root mean square error (RMSE); one-sample t-test; two-way analysis of variance; two-way analysis of variance without replication (water content level as a block) and the Friedman test; linear regression analysis.

Results. In the soils studied (content of physical clay, i.e., particles < 0.01 mm: group A 10–20%, B 40–50%, C 50–65%; electrical conductivity of the saturated-paste extract ECe = 0.28–0.61 dS/m; bulk density 1.35 and 1.52 g/cm³), the RMSE of the factory calibration (calculated from three-column means) was 0.006–0.014 m³/m³, below the adopted criterion of 0.02 m³/m³. The largest error among the ten analyzed water content levels (+0.027 m³/m³, three-column mean; group C at a bulk density of 1.35 g/cm³ and θ = 0.40 m³/m³) exceeded the manufacturer's stated accuracy (±0.025 m³/m³) by 0.002 m³/m³, i.e., by less than the expanded uncertainty of the reference method (0.005 m³/m³). At two wetter states recorded for this treatment but not included in the analysis (θ = 0.44 and 0.48 m³/m³), Δθ was +0.031 and +0.038 m³/m³. RMSE increased with the content of physical clay: on average 0.0065, 0.0100 and 0.0130 m³/m³ in groups A, B and C, respectively. MB was positive in all treatments; in group A it did not differ significantly from zero (+0.002 to +0.004 m³/m³; p = 0.066 and 0.32), whereas in groups B and C the sensor overestimated water content by 0.006–0.010 m³/m³ on average (descriptive one-sample t-tests over the ten water content levels, p = 0.014–0.035; the systematic component of the reference-method uncertainty (U = 0.005 m³/m³) was not included; the distinction between groups A and B at α = 0.05 is sensitive to the rounding of MB and RMSE). Treatment-level MB increased with maximum hygroscopic water content (linear fit over six treatments: 0.00066 m³/m³ per 1%; R² = 0.88; p = 0.006; p = 0.07 for the three group means) but was at least as closely related to ECe (R² = 0.96), so the contributions of pore-solution conductivity and surface conductivity could not be separated. At a bulk density of 1.35 g/cm³, over θ = 0.20–0.40 m³/m³, the error increased by 0.008, 0.014 and 0.021 m³/m³ in groups A, B and C, respectively. In the two-way analysis of variance (its residual term includes the dependence of the error on θ), differences between treatments explained about 10% of the total variability, and the texture effect (9%) was not significant (F(2, 54) = 2.74; p = 0.074). However, when the groups were compared at the same six θ levels (bulk density 1.35 g/cm³), the error increased in the order A < B ≤ C at every level (texture effect in a two-way analysis of variance without replication, descriptive: F(2, 10) = 6.92; p = 0.013). The bulk density effect was small and not significant (difference in MB −0.001 to −0.002 m³/m³; p = 0.43), whereas the complex refractive index model (CRIM) predicts an increase in the reading of about 0.009–0.010 m³/m³; no texture × bulk density interaction was detected (p = 0.98). At a bulk density of 1.52 g/cm³, the group C soil swelled (volumetric strain 8.3%; maximum water content θmax = 0.460 m³/m³ versus an initial porosity of 0.415 m³/m³); when the reference water content was calculated from the actual column volume, no additional error attributable to swelling was detected (in this comparison, the effect of swelling cannot be separated from those of bulk density and ECe). The CS616 factory calibration is suitable for the non-saline soils studied (10–20% and 40–65% physical clay) when an RMSE of 0.02 m³/m³ is acceptable, but in the studied soils with 40–65% physical clay it overestimated water content by 0.006–0.010 m³/m³ on average (up to 0.027 m³/m³ at θ = 0.40 m³/m³); soil-specific calibration or periodic gravimetric verification is therefore advisable.

Keywords: CS616 sensor, volumetric soil water content, factory calibration, systematic error, soil texture, bulk density, soil electrical conductivity, swelling, analysis of variance, chernozem

INTRODUCTION

Accurate measurement of volumetric soil water content (θ, m³/m³) is a prerequisite for optimizing irrigation scheduling and for precision agriculture [1, 2]. CS616 sensors (Campbell Scientific, USA) are water content reflectometers: the probe rods act as a waveguide connected to a bistable multivibrator, which switches when the reflected pulse returns, so that the output period τ (μs) depends on the travel time of the signal along the rods and, hence, on the effective dielectric permittivity εeff of the soil. With the rods in air, the oscillation frequency is about 70 MHz [3]. In mineral soils with low electrical conductivity, εeff is determined mainly by the volumetric water content [4, 5]. The manufacturer's standard (factory) calibration θ = f(τ) is a quadratic polynomial. Its accuracy specification (±0.025 m³/m³) is based on laboratory measurements in soils that were typically sandy loam and coarser. According to the manufacturer, the standard coefficients should be accurate in mineral soils with a bulk electrical conductivity below 0.5 dS/m, a bulk density below 1.55 g/cm³ and a clay (< 0.002 mm) content below 30%, and a soil-specific calibration is recommended when, among other conditions, the smectite clay content exceeds 25% [3]. The conductivity limit refers to the bulk soil conductivity, which is not equivalent to the electrical conductivity of the saturated-paste extract (ECe) used in this paper.

The applicability of the factory calibration is therefore limited by soil texture. In fine-textured soils, the bulk electrical conductivity comprises conduction through the pore solution and surface conduction in the electrical double layer of clay minerals, whose specific surface area is large. As θ increases, the liquid phase becomes continuous and the bulk electrical conductivity increases [5]; the resulting dielectric losses lengthen the measured period, so that the sensor may overestimate θ relative to the gravimetric reference (positive error), increasingly so in wet soil. Deviations of manufacturers' calibrations of dielectric sensors, including the CS616, from measured water content in soils of different texture and electrical conductivity have been reported in the USA [6–8]; however, no systematic data have been published for the chernozems and meadow-chernozem soils of the Krasnodar Krai, a viticultural region with predominantly heavy-textured soils [9]. In the soils studied here, the content of physical clay (particles < 0.01 mm according to N. A. Kachinskii [10]) reached 65%.

Soil bulk density ρb (g/cm³) changes the volume fractions of the solid, liquid and gaseous phases [11] and therefore εeff. The three-phase complex refractive index model (CRIM) [12], with permittivities of 80 for water, 4.7 for the solid phase and 1 for air, gives an apparent change in θ of (Δρb/ρs)·(√4.7 − 1)/(√80 − 1). For ρs = 2.60–2.67 g/cm³, this is an apparent increase in θ of about 0.0055–0.0057 m³/m³ per 0.1 g/cm³ increase in ρb at constant θ, because solids replace part of the air-filled pore space. The manufacturer also notes that, for some soils with a bulk density above 1.5 g/cm³, the slope of the sensor response (θ versus τ) decreases with increasing bulk density, and that the standard equations overestimate θ in soils of higher bulk density [3]. The density effect is expected to be secondary to the texture effect, but it can be tested under strictly controlled laboratory conditions.

Swelling is another factor that affects the accuracy of θ determination. Swelling of clay soils is caused by the disjoining pressure of diffuse double layers and increases with the initial dry density and the smectite content [13]; in dense heavy-textured soils it can therefore be considerable. Illite and the smectite mineral nontronite have been identified in vineyard soils (southern chernozems) of the Anapa district of the Krasnodar Krai [14]. The influence of swelling on CS616 readings has not been studied for the soils of the region.

Hypothesis. The error of the factory calibration, Δθ = θCS616 − θgrav, is determined mainly by the content of physical clay (through surface conductivity) and increases nonlinearly as θ approaches saturation; the effect of ρb is significant, and its magnitude depends on texture (interaction of the factors).

The aim of the study was to evaluate experimentally the mean bias (MB) and root mean square error (RMSE) of the CS616 factory calibration in non-saline soils of three texture groups at two bulk density levels and to partition the variability of the error among texture, bulk density and their interaction. The effect of salinity (ECe 1–4 dS/m) is considered in a separate paper by the authors.

MATERIALS AND METHODS

A. Experimental design.

A full factorial 3 × 2² design was used [15] (12 treatments: three texture groups × two ρb levels × two salinity levels). Stratified randomization was applied: the order of column measurements was randomized independently within each factor level [15]. Instrument drift was controlled by checking the equipment against standard solutions before each measurement series. This paper reports the six non-saline treatments (ECe < 1 dS/m, a conservative criterion adopted by the authors), i.e., a two-factor 3 × 2 subset of the design. The design is presented in Table I. Treatments are coded by texture group (A, B, C), bulk density level (1 for 1.35 g/cm³ and 2 for 1.52 g/cm³) and salinity level (n for non-saline); e.g., C2-n is group C at 1.52 g/cm³.

Table I – Full factorial 3 × 2² experiment and study conditions

| Factor | Symbol | Levels | Values | Rationale |
| --- | --- | --- | --- | --- |
| Texture | T | 3 | Group A: 10–20% physical clay (sandy loam); group B: 40–50% (medium to heavy loam); group C: 50–65% (heavy loam to light clay)* | Representative soils of the Krasnodar Krai [9] |
| Bulk density | ρb | 2 | 1.35 and 1.52 g/cm³ | Typical agrogenic range [9] |
| Salinity | ECe | 2 | Non-saline (< 1 dS/m), this paper; ECe 1–4 dS/m, separate publication | Background and elevated ranges below the saline-soil threshold of 4 dS/m [16] |
| Replication | – | – | n = 5 Kachinskii rings (undisturbed field samples); n = 3 for laboratory determinations of soil properties (Table II) and for repacked columns (calibration experiment) | Two independent replication types |
| θ levels | – | – | 10 analyzed equilibrium states per column (at least 10 recorded), from 0.40–0.50 m³/m³ (the upper limit differs between treatments; Results D) to air-dry (0.01–0.03 m³/m³) | Full range of water content |
| Total | – | 12 | 6 non-saline treatments × 10 levels × 3 columns = 180 column-level values; for analysis of variance (ANOVA), N = 60 values averaged over three columns | Balanced design |

Note: * class names according to the Kachinskii scale for soils of the steppe type of soil formation, which is used for chernozems and meadow-chernozem soils [10]; in the text, the groups are referred to by letter and by the content of physical clay (< 0.01 mm).

B. Scope and soils studied.

The results were obtained within the following ranges: physical clay 10–20% and 40–65%; ρb 1.35–1.52 g/cm³; ECe 0.28–0.61 dS/m; organic carbon content Corg < 3%; T = 20.0 ± 0.5 °C. Samples were collected in October 2025 in vineyard agroecosystems of the Krasnodar Krai from the 0–25 cm layer (in the vicinity of 45.33° N, 37.23° E). According to the soil classification of the USSR [17], the group A soil is an ordinary chernozem, the group B soil is a typical chernozem, and the group C soil is a meadow-chernozem soil. Composite samples were prepared from individual samples (increments) taken by the "envelope" method (GOST 17.4.4.02-2017 [18]; general sampling requirements according to GOST 17.4.3.01-2017 [19]); the sampling plot area was ≥ 1 ha, and the distance between sampling points was ≥ 50 m. Undisturbed samples (n = 5) were taken with 100 cm³ stainless-steel Kachinskii rings (GOST 5180-2015 [20]). Samples were transported at 4 ± 2 °C, and processing started within 24 h.

Samples were prepared at an air temperature of 20 ± 2 °C and a relative humidity of 50 ± 5%, following the principles of ISO/IEC 17025 [21]. Disturbed samples were dried at 40 ± 2 °C to constant mass; this temperature avoids alteration of clay minerals and dehydration of hydrated salts (e.g., gypsum), which can occur at higher temperatures [22]. The samples were passed through a sieve with a nominal aperture of 2 mm (GOST R 51568-99 [23]); the coarse fraction (> 2 mm) did not exceed 0.5% of the sample mass.

C. Analytical methods.

Particle-size distribution was determined by the Kachinskii pipette method [10] with dispersion of the sample by sodium pyrophosphate (Na4P2O7) [24]; the fractions were separated at the Kachinskii boundaries (error ≤ 3%; GOST 12536-2014 [25]). Physical clay (< 0.01 mm) includes, besides the clay fraction (< 0.001 mm), the medium and fine silt fractions (0.01–0.001 mm); its content is substantially higher than the content of clay (< 0.002 mm) used in the USDA system, and the ratio between them is soil-specific. The class names of the two systems are therefore not equivalent (e.g., Kachinskii "light clay" is not USDA "clay"). The particle density ρs was determined by pycnometry (GOST 5180-2015 [20]; expanded uncertainty U = 0.01 g/cm³, coverage factor k = 2). Maximum hygroscopic water content (MH) was determined by equilibration over a saturated K2SO4 solution (GOST 28268-89 [26]; permissible relative discrepancy between parallel determinations 7% at MH ≤ 5% and 5% at MH > 5%). The electrical conductivity of 1:5 soil–water extracts (σ1:5) was measured with an Expert-001 analyzer according to GOST 26423-85 [27] (permissible error 7.5% at σ1:5 ≤ 0.3 dS/m); ECe was measured directly in saturated-paste extracts [16]. For the six treatments in Table II, the paired values gave ECe/σ1:5 = 7.76–7.95 (mean 7.85; CV ≈ 1%). This near-constancy across the texture groups differs from the usual decrease of the ratio with increasing saturation percentage (and clay content) [28]. However, the classification of all samples as non-saline does not depend on the exact value of the ratio: since σ1:5 ≤ 0.044 dS/m in group A and ≤ 0.078 dS/m in groups B and C, ECe would remain below 1 dS/m for any ratio up to 22 in group A and up to 12 in groups B and C.

D. CS616 calibration experiment.

Repacked soil columns were prepared in polycarbonate cylinders (inner diameter 152.4 mm, height 300 mm; V = 5472 cm³) by compaction in three layers to the target bulk density (the achieved values are given in Table II). The probe (rods 300 mm long and 3.2 mm in diameter, spaced 32 mm apart [3]) was installed vertically along the column axis. The distance from the rod surfaces to the column wall was 59–73 mm, i.e., greater than the 50 mm zone around the rods that, according to the manufacturer, must contain only soil, but less than the clearance of at least 100 mm recommended for calibration containers [3]. Because the rod length equals the column height, the rod tips were at or near the column base, so the 50 mm soil-only condition was not met there. The possible influence of the column wall and base on τ was not quantified (see Limitations).

The temperature was maintained at 20.0 ± 0.5 °C (fluctuations in the chamber ±0.3 °C). According to the manufacturer's temperature correction [3], the temperature sensitivity of τ increases with water content: a deviation of 0.5 °C corresponds to ≤ 0.001 m³/m³ at θ ≤ 0.20 m³/m³ and to 0.004–0.005 m³/m³ at θ = 0.40–0.46 m³/m³. Insofar as the temperature fluctuated around the 20 °C reference of this correction, this component is random and does not bias MB; a constant offset within the ±0.5 °C tolerance cannot be excluded and would shift Δθ by at most the amounts given above. Initial saturation was achieved by capillary rise from below (48 h). Water content levels were created by stepwise isothermal drying. Equilibrium was assumed when the hourly change in the mean τ was < 0.025% and the hourly change in column mass was < 0.01%; equilibration took 8–36 h. At least ten equilibrium states were recorded for each column, and ten states per column were used in the analysis. For C1-n, two wetter states (θ = 0.44 and 0.48 m³/m³) were recorded but not included in the analysis (Results D).

The reference water content θgrav was determined by two methods: by weighing the whole column (resolution 0.1 g) and directly from ten micro-cores (12 mm in diameter, V = 15–17 cm³, taken within 35 mm of the probe axis and dried at 105 ± 2 °C according to GOST 5180-2015 [20]). The difference between the two methods did not exceed 0.008 m³/m³ (mean 0.003 ± 0.002 m³/m³); the expanded uncertainty of θgrav, including the contributions of weighing, column geometry and column heterogeneity, was U = 0.005 m³/m³ (k = 2). For the swelling soil of treatment C2-n, θgrav was calculated from the actual column volume V(θ) = π·(7.62 cm)²·(30.0 cm + ΔH(θ)), where ΔH is the rise of the soil surface; at θmax = 0.460 m³/m³, ΔH ≈ 25 mm and V ≈ 5930 cm³.

E. Statistical methods.

Sensor error: Δθ = θCS616 − θgrav. For each treatment (n = 10 water content levels; each value is the mean of three columns), the following were calculated:

MB = (1/n) Σ Δθi                                                  (1)

RMSE = √[(1/n) Σ (Δθi)²]                                           (2)

The standard error was calculated as SE(MB) = s/√n, where s is the sample standard deviation of Δθ (divisor n − 1), and the 95% confidence interval (CI) as MB ± t0.975;9·SE(MB), with t0.975;9 = 2.262; the null hypothesis MB = 0 was tested with a two-sided one-sample t-test (df = 9). Because the ten levels are successive states of the same columns and Δθ depends systematically on θ, they are not independent replicates; the CIs and p-values for MB are therefore descriptive and characterize the mean error over the tested θ range (the experimental unit is the column, n = 3 per treatment). To characterize the asymmetry of the error distribution, the index Jnl = [MB − (Δθmax + Δθmin)/2]/σΔθ was used, where σΔθ is the standard deviation of Δθ (divisor n). Jnl < 0 means that the mean lies below the midpoint of the range, which is consistent with a right-skewed distribution, in which a few large positive values form the upper tail. The acceptability criterion adopted by the authors was RMSE ≤ 0.02 m³/m³: for effective irrigation management, the accuracy of water content estimates should be of the order of 0.01–0.02 m³/m³ [29], and this RMSE threshold has been used to evaluate soil water sensors for irrigation scheduling [30].

Two-way ANOVA (texture × bulk density) was applied to the N = 60 column-averaged values of Δθ (6 treatments × 10 water content levels). Because the design is balanced, Type I, II and III sums of squares (SS) coincide. The within-treatment term contains both random error and the systematic dependence of Δθ on θ. The ANOVA therefore quantifies the share of the total variability of Δθ that is due to differences between treatments, and its F-tests for the between-treatment effects are conservative with respect to the θ-dependence. However, column-specific deviations that are constant across water content levels (e.g., probe-to-probe differences or errors in column volume) are not represented in this term, so these p-values, like those for MB, are descriptive (Limitation 7). To compare the texture groups at the same water content, Δθ at six selected θ levels at ρb = 1.35 g/cm³ (Results D) was analyzed by two-way ANOVA without replication, with the θ level as a block, and by the Friedman test (corrected for ties). Because these levels are successive states of the same columns and each texture group is represented here by a single treatment, these tests are also descriptive: they show whether the ranking of the three treatments is consistent across water contents, but their error term does not include between-column variability. In addition, simple and multiple linear regressions (n = 6) were used to relate treatment MB and RMSE to MH, which characterizes the specific surface area of the soil [31, 32], and MB to ECe and bulk density (indicator variable D = 0 for 1.35 and D = 1 for 1.52 g/cm³); the association between ECe and MH was characterized by Pearson's correlation coefficient. Differences in particle density among the texture groups were tested by one-way ANOVA. The sums of squares of the two-way ANOVA were reconstructed from the treatment statistics MB and RMSE (see Results C); therefore, residual diagnostics (normality and homogeneity of variance) could not be performed for this analysis, and the Friedman test, which does not assume normality, served as a robustness check for the texture effect. The significance level was α = 0.05. The treatment statistics (MB, RMSE, Δθmin, Δθmax and Jnl) were calculated by the authors in R (version 4.3.2) [33]; the confidence intervals, p-values, ANOVA, regressions, Friedman test and one-way ANOVA of particle density, as well as the rounding simulations, were calculated from the treatment-level values reported in Results A, B and D in Python 3.13 with the SciPy library (version 1.18) [34].

RESULTS

A. Physical and chemical properties of the samples.

All samples had ECe = 0.28–0.61 dS/m (Table II), i.e., below the adopted criterion of 1 dS/m and lower than the saline-soil threshold of 4 dS/m [16] by a factor of 6.6–14. The particle density decreased from group A (2.67 g/cm³) to group C (2.60 g/cm³; one-way ANOVA, p < 0.01), which may reflect a higher organic matter content in the finer-textured soils, since organic matter has a much lower particle density than soil minerals [11]. MH increased monotonically from group A to group C (A: 2.1–2.3%; B: 6.5–6.8%; C: 11.1–11.4%), consistent with the texture grouping and with reference ranges [31, 32]. Although σ1:5 and ECe were measured on disturbed material, in all three groups they were lower for the denser treatment (σ1:5 by 0.008–0.016 dS/m; ECe by 0.07–0.12 dS/m); the comparison of the two density levels is therefore partly confounded with salinity (Discussion C).

Table II – Physical and chemical properties of the non-saline treatments (n = 3, mean ± SD)

| Treatment | Texture group / physical clay, % | ρs, g/cm³ | MH, % | ρb, g/cm³ | σ1:5, dS/m | ECe, dS/m |
| --- | --- | --- | --- | --- | --- | --- |
| A1-n | A (10–20) | 2.67 ± 0.01 | 2.1 ± 0.2 | 1.35 ± 0.02 | 0.044 ± 0.003 | 0.35 ± 0.02 |
| A2-n | A (10–20) | 2.67 ± 0.01 | 2.3 ± 0.2 | 1.52 ± 0.02 | 0.036 ± 0.002 | 0.28 ± 0.02 |
| B1-n | B (40–50) | 2.63 ± 0.01 | 6.5 ± 0.3 | 1.35 ± 0.03 | 0.066 ± 0.004 | 0.52 ± 0.03 |
| B2-n | B (40–50) | 2.63 ± 0.01 | 6.8 ± 0.4 | 1.52 ± 0.03 | 0.058 ± 0.004 | 0.45 ± 0.03 |
| C1-n | C (50–65) | 2.60 ± 0.02 | 11.4 ± 0.6 | 1.35 ± 0.03 | 0.078 ± 0.005 | 0.61 ± 0.04 |
| C2-n† | C (50–65) | 2.60 ± 0.02 | 11.1 ± 0.6 | 1.52 ± 0.03 | 0.062 ± 0.004 | 0.49 ± 0.03 |

Note: treatment codes – see Methods A; ρs – particle density [20]; MH – maximum hygroscopic water content [26]; ρb – bulk density of the repacked columns; σ1:5 – electrical conductivity of the 1:5 extract [27]; ECe – electrical conductivity of the saturated-paste extract [16]. † The soil of treatment C2-n swelled during wetting, at θ > 0.40 m³/m³ (Results D).

B. Repeatability and mean error.

At each equilibrium state, five readings of τ were taken at 10 s intervals and averaged. The coefficient of variation of τ was < 0.14%, which corresponds to a standard deviation of a single reading ≤ 0.002 m³/m³ (≤ 0.0009 m³/m³ for the mean of five readings) over θ ≤ 0.50 m³/m³, i.e., less than the standard uncertainty of θgrav (u = U/2 = 0.0025 m³/m³). At the wettest levels, chamber temperature fluctuations (±0.3 °C) add a random component of up to about ±0.003 m³/m³. These components increase the within-treatment variance but do not bias MB; the analysis below concerns the systematic error Δθ.

θCS616 was calculated by substituting the measured τ into the manufacturer's standard equation θ = −0.0663 − 0.0063τ + 0.0007τ² (τ in μs) [3], and Δθ was obtained as θCS616 − θgrav. The statistics of Δθ are given in Table III.

Table III – Statistics of the error Δθ, 95% confidence intervals and asymmetry index Jnl (n = 10 water content levels, means of three columns)

| Treatment | ECe, dS/m | MB, m³/m³ | RMSE, m³/m³ | Δθmin / Δθmax, m³/m³ | 95% CI of MB, m³/m³ | p (MB = 0) | Jnl |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A1-n | 0.35 | +0.004 | 0.007 | −0.002 / +0.013 | [−0.0003, +0.0083] | 0.066 | −0.263 |
| A2-n | 0.28 | +0.002 | 0.006 | −0.004 / +0.011 | [−0.0023, +0.0063] | 0.32 | −0.268 |
| B1-n | 0.52 | +0.007 | 0.011 | −0.003 / +0.019 | [+0.0006, +0.0134] | 0.035 | −0.118 |
| B2-n | 0.45 | +0.006 | 0.009 | −0.004 / +0.016 | [+0.0009, +0.0111] | 0.025 | 0.000 |
| C1-n | 0.61 | +0.010 | 0.014 | −0.001 / +0.027 | [+0.0026, +0.0174] | 0.014 | −0.306 |
| C2-n | 0.49 | +0.008 | 0.012 | −0.002 / +0.022 | [+0.0013, +0.0147] | 0.025 | −0.225 |

Note: MB – mean bias; RMSE – root mean square error; 95% CI = MB ± 2.262·SE(MB), SE = s/√n with s = √[n/(n − 1)·(RMSE² − MB²)] calculated from the tabulated (rounded) MB and RMSE; p – two-sided one-sample t-test (df = 9); CIs and p are descriptive (Methods E); Jnl is sensitive to the rounding of MB, RMSE, Δθmin and Δθmax: recalculation from the tabulated values differs by ≤ 0.003 from the Jnl values shown.

In all six treatments, RMSE was 0.006–0.014 m³/m³ (Table III), below the acceptability criterion of 0.02 m³/m³ (Methods E). At all ten analyzed levels, the level-wise errors (three-column means) were within the manufacturer's stated accuracy of ±0.025 m³/m³ [3] except in C1-n at θ = 0.40 m³/m³, where Δθ (+0.027 m³/m³; Results D) exceeded it by 0.002 m³/m³, i.e., by less than the expanded uncertainty of θgrav (U = 0.005 m³/m³). RMSE increased with the content of physical clay: on average 0.0065, 0.0100 and 0.0130 m³/m³ in groups A, B and C, respectively (ratio about 1 : 1.5 : 2); MB followed the same trend. Point estimates of MB were positive in all treatments. In both group A treatments, the 95% CI of MB includes zero (p = 0.066 and 0.32). In all group B and C treatments, MB is positive and differs significantly from zero (+0.006 to +0.010 m³/m³; p = 0.014–0.035); these soils have MH ≥ 6.5%, which indicates a well-developed sorption surface [31, 32]. The conclusions for A1-n (p = 0.066) and B1-n (p = 0.035) are sensitive to the rounding of MB and RMSE (in simulations of the rounding error, p < 0.05 in 28% and 85% of cases, respectively), so the distinction between groups A and B at α = 0.05 is not robust (see also Limitation 8). The Jnl values were negative (−0.118 to −0.306) except in B2-n (0.000); the negative values are consistent with right-skewed distributions, in which the largest positive errors occur at high water contents (Results D). Jnl does not vary monotonically with texture.

C. Partitioning of the variability of the error.

The results of the two-way ANOVA are presented in Table IV.

Table IV – Two-way ANOVA of Δθ (N = 60; Type I SS = Type III SS)

| Source of variation | SS, (m³/m³)² | df | MS, (m³/m³)² | F | p | η² |
| --- | --- | --- | --- | --- | --- | --- |
| Texture (T) | 3.63·10⁻⁴ | 2 | 1.82·10⁻⁴ | 2.74 | 0.074 | 0.091 |
| Bulk density (ρb) | 4.17·10⁻⁵ | 1 | 4.17·10⁻⁵ | 0.63 | 0.43 | 0.010 |
| Interaction T × ρb | 3.3·10⁻⁶ | 2 | 1.7·10⁻⁶ | 0.03 | 0.98 | 0.001 |
| Within treatments | 3.58·10⁻³ | 54 | 6.63·10⁻⁵ | – | – | 0.898 |
| Total | 3.99·10⁻³ | 59 | – | – | – | 1.000 |

Note: SS – sum of squares; MS – mean square; η² = SSsource/SStotal; α = 0.05. SS were calculated from the treatment MB and RMSE (Table III): SSwithin = n·Σ(RMSE² − MB²) and SStotal = n·ΣRMSE² − N·MB̄² (sums over the six treatments; n = 10; N = 60; MB̄ – mean of the six MB values). Because MB and RMSE are rounded to 0.001 m³/m³, F and p are approximate: in 20,000 simulations of the rounding error (uniform ±0.0005 m³/m³), the central 95% range of p for texture was 0.04–0.12 and that of η² was 0.07–0.11. The within-treatment term includes the systematic dependence of Δθ on θ. The 95% CI of the bulk density effect (mean MB at 1.52 minus mean MB at 1.35 g/cm³ = −0.0017 m³/m³) was calculated as −0.0017 ± t0.975;54·√(2·MSwithin/30), t0.975;54 = 2.005, i.e., −0.006 to +0.003 m³/m³.

Differences between treatments account for about 10% of the total variability of Δθ (texture 9.1%, bulk density 1.0%, interaction 0.1%); the remaining 90% is within-treatment variability, which is associated mainly with the increase in Δθ with θ (Table V). Tested against this within-treatment variance, the texture effect is not significant at α = 0.05 (F = 2.74; p = 0.074), and the effects of bulk density (p = 0.43) and of the interaction (p = 0.98) are not significant either. Because the error term includes the θ-dependence, this test is conservative for texture in this respect, although column-level variability is not represented in its error term (Methods E). When the texture groups are compared at the same water content (Table V, ρb = 1.35 g/cm³, θ level as a block), Δθ increases in the order A < B ≤ C at each of the six levels, and the texture effect is significant in the descriptive sense defined in Methods E: F(2, 10) = 6.92, p = 0.013; Friedman test (corrected for ties) χ²(2) = 11.27, p = 0.004. At the level of treatment means, MB increased linearly with MH: MB = 0.0017 + 0.00066·MH (R² = 0.88; p = 0.006; because the two density treatments of each group were prepared from the same soil, the regression on the three group means gives p = 0.07); the corresponding relationship for RMSE is RMSE = 0.0050 + 0.00072·MH (R² = 0.90; p = 0.004). With bulk density added as an indicator variable, MB = 0.0026 + 0.00066·MH − 0.0017·D (R² = 0.98; p = 0.02 for D; n = 6); the coefficient of D is stable under the rounding of the published values (−0.0022 to −0.0013), but its p-value is not (p < 0.05 in 76% of the rounding simulations). With ECe as the only predictor, MB = −0.0044 + 0.0235·ECe (R² = 0.96; p < 0.001); with ECe instead of MH in the model with D, MB = −0.0050 + 0.0243·ECe + 0.0004·D (R² = 0.97; p = 0.51 for D); with both MH and ECe as predictors, only the ECe coefficient was significant (p = 0.017; p = 0.11 for MH; this pattern held in about two thirds of the rounding simulations). Thus, the mean level of the error is associated with texture (MH) and with ECe, whose effects cannot be separated here, whereas water content governs its variation within a treatment.

The differences in MB between the density levels (MB at 1.52 minus MB at 1.35 g/cm³) were −0.002 m³/m³ in group A, −0.001 in group B and −0.002 in group C, i.e., of the same sign and practically the same magnitude in all groups, and negligible compared with the 0.02 m³/m³ criterion. No interaction between texture and bulk density was detected in the range of 1.35–1.52 g/cm³.

D. Nonlinear dependence and swelling.

To illustrate the nonlinearity, Table V presents Δθ at selected levels of θ for the three texture groups at ρb = 1.35 g/cm³.

Table V – Dependence of the error Δθ (m³/m³) on volumetric water content θ for the three texture groups at ρb = 1.35 g/cm³

| θ, m³/m³ | 0.05 | 0.10 | 0.20 | 0.30 | 0.38 | 0.40 |
| --- | --- | --- | --- | --- | --- | --- |
| A1-n | 0.000 | 0.001 | 0.003 | 0.005 | 0.010 | 0.011 |
| B1-n | 0.001 | 0.002 | 0.005 | 0.009 | 0.017 | 0.019 |
| C1-n | 0.001 | 0.002 | 0.006 | 0.012 | 0.022 | 0.027 |

Note: means of n = 3 columns at six selected water content levels. For C1-n, two wetter equilibrium states (θ = 0.44 and 0.48 m³/m³; Δθ = 0.031 and 0.038 m³/m³) were also recorded; they are not among the ten levels analyzed in Tables III and IV, i.e., near saturation the error in group C was larger than at the levels shown (up to +0.038 m³/m³).

At ρb = 1.35 g/cm³, over θ = 0.20–0.40 m³/m³, Δθ increased by 0.008 m³/m³ in A1-n, by 0.014 m³/m³ in B1-n and by 0.021 m³/m³ in C1-n (2.6 times that in A1-n). The increase is thus monotonic in the content of physical clay but less than proportional to MH: MH differs between C1-n and A1-n by a factor of 5.4 (the mid-range contents of physical clay in groups C and A differ by a factor of 3.8, but because only the group ranges are reported, the actual ratio may lie between 2.5 and 6.5). The nonlinear increase is consistent with the increasing continuity of the liquid phase and the increasing bulk electrical conductivity in wet soil [5].

In treatment C2-n (ρb = 1.52 g/cm³), the maximum water content θmax = 0.460 m³/m³ was recorded; it exceeds the initial total porosity φ0 = 1 − ρb/ρs = 0.415 m³/m³ by 0.045 m³/m³, which indicates that the soil swelled. At θmax, the rise of the soil surface was ΔH ≈ 25 mm, i.e., the vertical swelling strain of the laterally confined column, which equals its volumetric strain, was εv ≈ 8.3%. With this strain, the porosity of the swollen soil, 1 − (1 − φ0)/(1 + εv) = 1 − 0.585/1.083 = 0.460 m³/m³, equals θmax; that is, θmax corresponds to a water-saturated swollen soil. In magnitude, such a strain corresponds to a medium-swelling soil according to the Russian classification by free swelling strain (0.08 < εsw ≤ 0.12; GOST 25100-2020 [35]) and to a "high" swelling potential according to Seed et al. (5–25% swell) [36]; both classifications refer to standardized test conditions (initial water content, compaction state, surcharge) that differ from those of this experiment, so the comparison is indicative only. With θgrav calculated from the actual column volume, no additional systematic error attributable to swelling was detected at the treatment level: the MB of C2-n (+0.008 m³/m³) does not exceed that of C1-n (+0.010 m³/m³); this comparison is, however, confounded by the differences in bulk density and ECe. The comparison is also geometrically approximate: at the wettest levels the swollen soil (soil height ≈ 325 mm) extended beyond the 300 mm rods, so the probe did not sample the whole soil volume used for the reference value, and V(θ) (Methods D) accounts only for vertical deformation, not for possible lateral shrinkage (separation from the column wall) during subsequent drying. No measurable swelling was observed in the less dense treatment C1-n (ρb = 1.35 g/cm³).

DISCUSSION

A. Acceptability of the factory calibration for the non-saline soils studied.

The RMSE values of 0.006–0.014 m³/m³ obtained for all non-saline treatments (ECe = 0.28–0.61 dS/m) are below the criterion of 0.02 m³/m³. The factory calibration of CS616 sensors therefore meets this criterion in the three non-saline mineral soils studied (10–20% and 40–65% physical clay); no correction appears necessary in group A (MB +0.002 to +0.004 m³/m³), whereas one may be needed in groups B and C (MB +0.006 to +0.010 m³/m³; Recommendations 1 and 2), although the distinction between groups A and B at α = 0.05 is sensitive to rounding (Results B). These RMSE values refer to three-column means obtained under controlled laboratory conditions; they do not include the probe-to-probe variability stated by the manufacturer (±0.005 m³/m³ in dry and ±0.015 m³/m³ in saturated soil [3]), so the error of a single field installation can be larger. Both RMSE and MB increase with the content of physical clay, and in groups B and C the sensor overestimates θ by 0.006–0.010 m³/m³ on average and by up to 0.027 m³/m³ at the analyzed levels (C1-n at θ = 0.40 m³/m³; Table V).

B. Mechanism of the texture effect.

The specific surface area of smectites reaches 600–800 m²/g, whereas that of sand and silt particles is smaller by several orders of magnitude [37]. In fine-textured soils, surface conduction in the electrical double layer adds to the conduction through the pore solution, and both increase as the liquid phase becomes continuous in wet soil [5]. The resulting dielectric losses attenuate the signal on the rods, so that the multivibrator switches later and the output period lengthens [3, 5]; in addition, at the relatively low operating frequency of the CS616, interfacial polarization in clay-rich soils increases the apparent permittivity [38]. Clay-bound water, whose permittivity is lower than that of free water, acts in the opposite direction (underestimation of θ) [3, 5]. The positive MB, which increased with MH and ECe, is therefore consistent with a predominance of conductivity and polarization effects in the soils studied, although boundary effects of the column (Limitation 4) and the systematic component of the uncertainty of θgrav (Limitation 8) may contribute to its absolute level. This mechanism is consistent with the RMSE in group C being twice that in group A and with the increase in the errors at high water contents. The linear dependence of MB and RMSE on MH (a measure of the specific surface area; R² = 0.88–0.90) is also consistent with it. However, MB is explained by ECe at least as well (R² = 0.96; n = 6), and ECe and MH are correlated (r = 0.88; p = 0.021); in a two-predictor regression, only ECe retained a significant coefficient (in about two thirds of the rounding simulations). The data therefore do not discriminate between the contributions of pore-solution conductivity and surface conductivity. According to the manufacturer, the sensor response begins to change when the soil-solution electrical conductivity exceeds 2 dS/m [3]; the ECe values of the soils studied (0.28–0.61 dS/m) are several times lower, although the pore-solution conductivity at water contents below saturation exceeds ECe. The increase in Δθ over θ = 0.20–0.40 m³/m³ was 2.6 times as large in C1-n as in A1-n, whereas MH differs by a factor of 5.4; the error therefore increases less than proportionally to MH (for physical clay, only group ranges are available, and their ratio, 2.5–6.5 with a mid-range value of 3.8, does not allow a firm conclusion).

C. Role of bulk density.

An increase in ρb from 1.35 to 1.52 g/cm³ reduces total porosity by 0.064–0.065 m³/m³. At constant θ, solids (ε ≈ 4.7) replace air (ε = 1), so εeff and the sensor reading should increase; CRIM [12] predicts an apparent increase in θ of about +0.009 to +0.010 m³/m³ (0.0094–0.0096 m³/m³, depending on ρs). The observed differences in MB between the density levels were, however, small and negative (−0.001 to −0.002 m³/m³): in each group, the difference was opposite in sign to the CRIM prediction and below it by about 0.010–0.012 m³/m³. The density effect estimated from the ANOVA was −0.0017 m³/m³ (95% CI −0.006 to +0.003 m³/m³; descriptive, see Methods E), and this interval also excludes the CRIM prediction, so the expected positive shift was not observed. The manufacturer's note that the slope of the sensor response decreases with increasing bulk density in some soils above 1.5 g/cm³ [3] does not explain this result. Because the manufacturer plots θ against τ, a smaller slope means a longer period at a given θ, and the manufacturer states that the standard equations overestimate θ in soils of higher bulk density [3], i.e., a shift of the same sign as that predicted by CRIM. A possible explanation is that in each texture group the ECe measured for the denser treatment was lower (by 0.07–0.12 dS/m), which reduces the conductivity-induced lengthening of τ and may offset the dielectric effect of the solid phase. Consistent with this, when ECe replaces MH in the regression, the coefficient of D becomes +0.0004 m³/m³ (p = 0.51). In addition, the comparison is not made at constant θ: MB is an average over treatment-specific θ ranges, whose upper limits differ between treatments (Table I; note to Table V), whereas the CRIM prediction refers to constant θ; because Δθ increases with θ, differences in the θ ranges can shift MB independently of any density effect. Within 1.35–1.52 g/cm³, the net effect of bulk density on the error was therefore small (about −0.002 m³/m³) compared with the texture effect, and no interaction between the factors was detected; however, a small real density effect of this magnitude can be neither confirmed nor ruled out.

D. Comparison with published data.

Direct comparison with published data is limited, because the studies differ in sensors, soils, texture classification and experimental design. Varble and Chávez [7] found that, under laboratory and field conditions, the factory calibrations of the CS616 and other sensors did not consistently achieve the required accuracy in the sandy clay loam, loamy sand and clay loam soils of eastern Colorado, and that the errors of the CS616 increased with bulk electrical conductivity when salt was added. Vaz et al. [8] evaluated eight sensors, including the CS616, in seven texturally different soils and found that the factory calibrations for mineral soils generally performed well, although some inconsistencies were identified; soil-specific calibration of the CS616 yielded an accuracy of about 0.025 m³/m³. Bosch [6] reported that capacitance probes (not the CS616) estimated θ in Coastal Plain soils within ±0.05 m³/m³ of the observed values in the laboratory, with greater variability in the field, and that soil-specific calibration improved accuracy. In a systematic review of soil water sensor evaluations in the USA, only 11% of the RMSE values obtained with factory calibrations met the 0.02 m³/m³ threshold, compared with 48% for site-specific calibrations [30]. Our RMSE values (0.006–0.014 m³/m³) are lower than most factory-calibration RMSE values reported in these studies, probably mainly because they were obtained under controlled laboratory conditions at equilibrium and refer to three-column means. The increase in the error with ECe observed here agrees qualitatively with the salt-addition results of Varble and Chávez [7]. Kelleners et al. [38] showed that the real permittivity of bentonite (a strongly dispersive smectitic clay), unlike that of fine sand, depends strongly on frequency below about 500 MHz, which affects the calibration of dielectric sensors operating in this range; this is consistent with the larger errors in groups B and C. Because physical clay according to Kachinskii and USDA clay (< 0.002 mm) are not equivalent, quantitative comparisons are approximate; differences in clay mineralogy, which was not determined in this study, may also contribute to the differences.

E. Confounding of texture and salinity; practical implications.

In the soils studied, ECe increased with the content of physical clay (r = 0.88 between ECe and MH), so the effects of salinity and texture on MB cannot be separated statistically in this experiment. Attributing MB to texture (physical clay ≥ 40%, MH ≥ 6.5%) rather than to salinity is only a working interpretation, based on physical considerations (Discussion B) and on the low ECe values (0.28–0.61 dS/m); in the regressions, ECe explains MB at least as well as MH does, so a contribution of salinity cannot be excluded and may be dominant. Because a positive error of 0.006–0.010 m³/m³ on average was found at ECe = 0.45–0.61 dS/m, i.e., far below the saline-soil threshold, a similar error may occur in other heavy-textured non-saline soils of the region; this should be verified on further soils.

RECOMMENDATIONS

1. For non-saline soils with 10–20% physical clay (group A) similar to that studied (Limitation 10), the factory calibration can be used without correction; soils with 20–40% physical clay were not tested.

2. For non-saline soils with 40–65% physical clay (MH ≥ 6.5%) similar to those studied (Limitation 10), the factory calibration meets the 0.02 m³/m³ RMSE criterion but overestimates θ by 0.006–0.010 m³/m³ on average and by up to 0.027 m³/m³ at θ = 0.40 m³/m³; closer to saturation, the overestimation was larger (up to 0.038 m³/m³ at θ = 0.48 m³/m³ in states not included in the analysis; note to Table V). Where this matters (e.g., irrigation thresholds near field capacity), soil-specific calibration or periodic gravimetric verification (1–2 samplings per irrigation season) is recommended. A constant offset based on MH is insufficient, because Δθ increases with θ and is associated with ECe as well as with MH.

3. Within the bulk density range of 1.35–1.52 g/cm³, the density effect is small compared with the texture effect, so corrections can be developed primarily for soil texture and salinity; in denser soils (ρb > 1.52 g/cm³, outside the range tested), the density effect should be verified.

4. In laboratory calibration of swelling soils (in this study, 50–65% physical clay at ρb = 1.52 g/cm³), the rise of the soil surface should be monitored during wetting and the reference water content calculated from the actual volume (in a laterally confined column, the vertical strain equals the volumetric strain; GOST 12248.6-2020 [39]).

Limitations. (1) The data refer to mineral soils with Corg < 3% at T = 20.0 ± 0.5 °C and do not apply to organic horizons. (2) The factory calibration was evaluated on the drying branch only; readings during wetting may differ, which was not tested. (3) According to the manufacturer's temperature correction [3], a deviation of 10 °C from 20 °C changes uncorrected readings by about 0.006–0.007 m³/m³ in dry soil and by 0.02–0.08 m³/m³ at θ ≈ 0.2–0.4 m³/m³, so field data require temperature correction. (4) The rod-to-wall distance (59–73 mm) is smaller than the 100 mm clearance recommended by the manufacturer for calibration containers [3], and the rod length (300 mm) equals the column height, so the column wall and base may have affected τ; these boundary effects were not quantified and may contribute to Δθ. In addition, the micro-cores for the direct reference method were taken within 35 mm of the probe axis, i.e., less than 40 mm from the rods, within the 50 mm zone to which the sensor is most sensitive and which, according to the manufacturer, must contain only soil [3]; the effect of the sampling holes on τ at subsequent (drier) levels was not quantified. (5) The bulk electrical conductivity was not measured, so compliance with the manufacturer's condition (< 0.5 dS/m) and the contribution of surface conductivity could not be verified directly. The clay fraction (< 0.001 mm by Kachinskii) is not equivalent to USDA clay (< 0.002 mm), to which the manufacturer's 30% limit refers. Because USDA clay is part of physical clay (< 0.01 mm), it cannot exceed 20% in group A, which therefore meets this limit, whereas compliance could not be assessed for groups B and C. (6) RMSE was computed from three-column means and does not include probe-to-probe variability; because the mean square error of a three-column mean cannot exceed the average mean square error of the individual columns, the reported RMSE values are lower bounds for the root mean square of the RMSEs of the individual columns; the RMSE of an individual sensor can be smaller or larger. The number of probes and their allocation to columns and treatments are not reported here. Because the probe-to-probe variability stated by the manufacturer (±0.005 m³/m³ in dry and ±0.015 m³/m³ in saturated soil [3]) is of the same order as the differences between the texture groups (0.0025–0.006 m³/m³ in group-mean MB, Table III; 0.008–0.016 m³/m³ in Δθ at θ = 0.40 m³/m³, Table V), a contribution of probe differences to these differences cannot be excluded unless the same probes were used in all treatments. (7) The ANOVA and the regressions were computed from treatment-level statistics. The between-treatment effects were tested against within-treatment variability that includes the θ-dependence but not the between-column variability, and the block ANOVA and Friedman test on Table V and the six-point regressions likewise ignore between-column variability (a column- or probe-specific offset cannot be distinguished from a treatment effect). The six points of the regressions and of the ECe–MH correlation represent only three soils at two bulk densities (with the three group means, the regression of MB on MH gives R² = 0.99 but p = 0.07, and r(ECe, MH) = 0.97 with p = 0.17; df = 1). Residual diagnostics could not be performed. The conclusions on texture, bulk density and their interaction should therefore be confirmed by a reanalysis of the column-level data (N = 180) with a mixed model (texture, bulk density and water content level as fixed effects; column as a random effect). (8) The CIs and p-values of MB reflect only the scatter of Δθ among water content levels; they do not include the systematic component of the uncertainty of θgrav (U = 0.005 m³/m³; mean difference between the two reference methods 0.003 m³/m³), which is comparable to MB in groups A and B. (9) Salinity (ECe 1–4 dS/m) is deliberately excluded from this paper. (10) Each texture group was represented by one soil, from which both density treatments were prepared (Table II), and the clay mineralogy of the samples was not determined; the texture effect is therefore confounded with the other properties of these soils (soil type, ECe, organic matter, mineralogy), and statements for ranges of physical clay are extrapolations from single soils.

CONCLUSIONS

1. In the non-saline mineral soils studied (ECe = 0.28–0.61 dS/m; physical clay 10–20% and 40–65%; bulk density 1.35 and 1.52 g/cm³), the RMSE of the CS616 factory calibration (three-column means, laboratory conditions) was 0.006–0.014 m³/m³, below the criterion of 0.02 m³/m³ in all treatments. The largest error among the ten analyzed water content levels (+0.027 m³/m³, C1-n at θ = 0.40 m³/m³; three-column mean) exceeded the manufacturer's stated accuracy (±0.025 m³/m³) by 0.002 m³/m³, i.e., by less than the expanded uncertainty of θgrav (0.005 m³/m³). At two wetter states recorded for this treatment but not included in the analysis (θ = 0.44 and 0.48 m³/m³), Δθ was +0.031 and +0.038 m³/m³ (note to Table V). The factory calibration is applicable to these soils when an RMSE of 0.02 m³/m³ is acceptable; for the soils studied with 40–65% physical clay, soil-specific calibration or periodic gravimetric verification is recommended.

2. Point estimates of MB were positive in all treatments and increased with the content of physical clay: MB did not differ significantly from zero in group A (p = 0.066 and 0.32) and was significantly greater than zero in groups B and C (+0.006 to +0.010 m³/m³; descriptive one-sample t-tests, p = 0.014–0.035); the distinction between groups A and B at α = 0.05 is, however, sensitive to the rounding of MB and RMSE (Results B). At the level of treatment means, MB increased with MH (linear fit over six treatments: 0.00066 m³/m³ per 1% MH; R² = 0.88; p = 0.006; p = 0.07 for the three texture-group means); because MB was at least as closely associated with ECe (R² = 0.96), the contributions of texture and salinity could not be separated.

3. Differences between treatments accounted for about 10% of the total variability of Δθ (texture 9.1%; F = 2.74; p = 0.074 in the ANOVA whose error term includes the θ-dependence); at six matched water contents (ρb = 1.35 g/cm³), Δθ increased in the order A < B ≤ C at every level (descriptive F(2, 10) = 6.92; p = 0.013). The error increased nonlinearly with θ, most strongly in group C (+0.021 m³/m³ over θ = 0.20–0.40 m³/m³ versus +0.008 m³/m³ in group A at ρb = 1.35 g/cm³). The effect of bulk density was small, negative in all three groups and not significant in the ANOVA (about −0.002 m³/m³; p = 0.43; 95% CI −0.006 to +0.003 m³/m³, which excludes the CRIM prediction of about +0.009 to +0.010 m³/m³); no texture × bulk density interaction was detected (p = 0.98). The hypothesis is thus only partly supported: the association of the error with texture and its nonlinear increase with θ (over the range shown in Table V, up to θ = 0.40 m³/m³) agree with it, but because the association with texture could not be separated from that with ECe, the proposed mechanism (surface conductivity) was not confirmed. The hypothesized significant, texture-dependent effect of bulk density was not confirmed either.

4. At ρb = 1.52 g/cm³, the group C soil swelled (εv = 8.3%), consistent with the excess of θmax over the initial porosity (0.045 m³/m³). When the actual column volume was taken into account, no additional error attributable to swelling was detected (MB of C2-n +0.008 m³/m³ versus +0.010 m³/m³ in C1-n; in this comparison, the effect of swelling cannot be separated from those of bulk density and ECe, and the comparison is insensitive to errors confined to the wettest levels). In laboratory calibration of swelling soils, monitoring of vertical deformation is recommended (GOST 12248.6-2020 [39]).

ACKNOWLEDGMENTS

The authors thank the staff of the Department of Hydraulics and Agricultural Water Supply of the Kuban State Agrarian University named after I. T. Trubilin for their assistance with the laboratory work and the discussion of the results, and the partner farms of the Krasnodar Krai for providing vineyard plots for soil sampling.

Funding. The study received no external funding.

Conflict of interest. The authors declare no conflict of interest.

БЛАГОДАРНОСТИ

Авторы благодарят сотрудников кафедры строительства и эксплуатации водохозяйственных объектов Кубанского государственного аграрного университета имени И. Т. Трубилина за помощь в проведении лабораторных исследований и обсуждении результатов, а также хозяйства-партнёры Краснодарского края за предоставление участков виноградников для отбора почвенных образцов.

Финансирование. Исследование не имело спонсорской поддержки.

Конфликт интересов. Авторы заявляют об отсутствии конфликта интересов.

ВКЛАД АВТОРОВ

Все авторы сделали эквивалентный вклад в подготовку публикации.

Вербицкий А. Ю. — концепция и дизайн исследования.

Приходько И. А. — сбор и статистическая обработка материала.

Болдырева Л. М. — анализ и интерпретация данных, написание текста.

Кесафоти Х. Е. — редактирование рукописи.

AUTHOR CONTRIBUTIONS

The authors contributed equally to this article.

Verbitsky A. Yu.: conceptualization and study design.

Prikhodko I. A.: data collection and statistical processing.

Boldyreva L. M.: data analysis and interpretation, writing of the manuscript.

Kesafoti Kh. E.: review and editing.

REFERENCES

1. Hillel D. Introduction to Environmental Soil Physics. Amsterdam: Elsevier Academic Press, 2004. 494 p.

2. Brady N.C., Weil R.R. Elements of the Nature and Properties of Soils. 3rd ed. Upper Saddle River, NJ: Pearson Prentice Hall, 2010. 614 p.

3. Campbell Scientific, Inc. CS616 and CS625 Water Content Reflectometers: Product Manual. Rev. 01/2026. Logan, UT: Campbell Scientific, Inc., 2026. Available at: https://s.campbellsci.com/documents/us/manuals/cs616.pdf (accessed 06.10.2026).

4. Topp G.C., Davis J.L., Annan A.P. Electromagnetic determination of soil water content: measurements in coaxial transmission lines. Water Resources Research. 1980;16(3):574–582. DOI: 10.1029/WR016i003p00574

5. Robinson D.A., Jones S.B., Wraith J.M., et al. A review of advances in dielectric and electrical conductivity measurement in soils using time domain reflectometry. Vadose Zone Journal. 2003;2(4):444–475. DOI: 10.2136/vzj2003.4440

6. Bosch D.D. Comparison of capacitance-based soil water probes in coastal plain soils. Vadose Zone Journal. 2004;3(4):1380–1389. DOI: 10.2136/vzj2004.1380

7. Varble J.L., Chávez J.L. Performance evaluation and calibration of soil water content and potential sensors for agricultural soils in eastern Colorado. Agricultural Water Management. 2011;101(1):93–106. DOI: 10.1016/j.agwat.2011.09.007

8. Vaz C.M.P., Jones S.B., Meding M., Tuller M. Evaluation of standard calibration functions for eight electromagnetic soil moisture sensors. Vadose Zone Journal. 2013;12(2):1–16. DOI: 10.2136/vzj2012.0160

9. Val'kov V.F., Shtompel' Yu.A., Trubilin I.T., et al. Pochvy Krasnodarskogo kraya, ikh ispol'zovanie i okhrana [Soils of the Krasnodar Krai, Their Use and Protection]. Rostov-on-Don: Izdatel'stvo SKNTs VSh, 1996. 192 p. (In Russ.)

10. Kachinskii N.A. Mekhanicheskii i mikroagregatnyi sostav pochvy, metody ego izucheniya [Mechanical and Microaggregate Composition of Soil and Methods of Its Study]. Moscow: Izdatel'stvo AN SSSR, 1958. 192 p. (In Russ.)

11. Shein E.V. Kurs fiziki pochv [A Course in Soil Physics]. Moscow: Izdatel'stvo MGU, 2005. 430 p. (In Russ.)

12. Roth K., Schulin R., Flühler H., Attinger W. Calibration of time domain reflectometry for water content measurement using a composite dielectric approach. Water Resources Research. 1990;26(10):2267–2273. DOI: 10.1029/WR026i010p02267

13. Nelson J.D., Miller D.J. Expansive Soils: Problems and Practice in Foundation and Pavement Engineering. New York: Wiley, 1992. 259 p.

14. Temerdashev Z., Abakumov A., Bolshov M., et al. Impact of soil profile mineralogy on the elemental composition of Chardonnay grapes and wines in the Anapa region. Beverages. 2026;12(1):1. DOI: 10.3390/beverages12010001

15. Adler Yu.P., Markova E.V., Granovskii Yu.V. Planirovanie eksperimenta pri poiske optimal'nykh uslovii [Design of Experiments in the Search for Optimal Conditions]. 2nd ed. Moscow: Nauka, 1976. 279 p. (In Russ.)

16. Richards L.A. (ed.) Diagnosis and Improvement of Saline and Alkali Soils. Agriculture Handbook No. 60. Washington, DC: USDA, 1954. 160 p.

17. Egorov V.V., Fridland V.M., Ivanova E.N., et al. Klassifikatsiya i diagnostika pochv SSSR [Classification and Diagnostics of Soils of the USSR]. Moscow: Kolos, 1977. 223 p. (In Russ.)

18. GOST 17.4.4.02-2017. Okhrana prirody. Pochvy. Metody otbora i podgotovki prob dlya khimicheskogo, bakteriologicheskogo, gel'mintologicheskogo analiza [Nature protection. Soils. Methods for sampling and preparation of soil for chemical, bacteriological, helmintological analysis]. Moscow: Standartinform, 2018. (In Russ.)

19. GOST 17.4.3.01-2017. Okhrana prirody. Pochvy. Obshchie trebovaniya k otboru prob [Nature protection. Soils. General requirements for sampling]. Moscow: Standartinform, 2018. (In Russ.)

20. GOST 5180-2015. Grunty. Metody laboratornogo opredeleniya fizicheskikh kharakteristik [Soils. Laboratory methods for determination of physical characteristics]. Moscow: Standartinform, 2016. (In Russ.)

21. ISO/IEC 17025:2017. General requirements for the competence of testing and calibration laboratories. Geneva: ISO, 2017. 30 p.

22. Mitchell J.K., Soga K. Fundamentals of Soil Behavior. 3rd ed. Hoboken, NJ: John Wiley & Sons, 2005. 577 p.

23. GOST R 51568-99 (ISO 3310-1-90). Sita laboratornye iz metallicheskoi provolochnoi setki. Tekhnicheskie usloviya [Test sieves of metal wire cloth. Specifications]. Moscow: Izdatel'stvo standartov, 2000. (In Russ.)

24. Ganzhara N.F., Borisov B.A., Baibekov R.F. Praktikum po pochvovedeniyu [Practical Course in Soil Science]. Ed. by N.F. Ganzhara. Moscow: Agrokonsalt, 2002. 280 p. (In Russ.)

25. GOST 12536-2014. Grunty. Metody laboratornogo opredeleniya granulometricheskogo (zernovogo) i mikroagregatnogo sostava [Soils. Methods of laboratory granulometric (grain-size) and microaggregate distribution]. Moscow: Standartinform, 2015. (In Russ.)

26. GOST 28268-89. Pochvy. Metody opredeleniya vlazhnosti, maksimal'noi gigroskopicheskoi vlazhnosti i vlazhnosti ustoichivogo zavyadaniya rastenii [Soils. Methods of determination of moisture, maximum hygroscopic moisture and moisture of steady plant fading]. Moscow: Izdatel'stvo standartov, 1989. (In Russ.)

27. GOST 26423-85. Pochvy. Metody opredeleniya udel'noi elektricheskoi provodimosti, pH i plotnogo ostatka vodnoi vytyazhki [Soils. Methods for determination of specific electric conductivity, pH and solid residue of water extract]. Moscow: Izdatel'stvo standartov, 1985. (In Russ.)

28. Slavich P.G., Petterson G.H. Estimating the electrical conductivity of saturated paste extracts from 1:5 soil:water suspensions and texture. Australian Journal of Soil Research. 1993;31(1):73–81. DOI: 10.1071/SR9930073

29. Hignett C., Evett S. Direct and surrogate measures of soil water content. In: Field Estimation of Soil Water Content: A Practical Guide to Methods, Instrumentation and Sensor Technology. Training Course Series No. 30. Vienna: IAEA, 2008. P. 1–22.

30. Datta S., Taghvaeian S. Soil water sensors for irrigation scheduling in the United States: a systematic review of literature. Agricultural Water Management. 2023;278:108148. DOI: 10.1016/j.agwat.2023.108148

31. Voronin A.D. Osnovy fiziki pochv [Fundamentals of Soil Physics]. Moscow: Izdatel'stvo MGU, 1986. 243 p. (In Russ.)

32. Kachinskii N.A. Fizika pochvy [Soil Physics]. Part 1. Moscow: Vysshaya shkola, 1965. 323 p. (In Russ.)

33. R Core Team. R: A language and environment for statistical computing. Vienna: R Foundation for Statistical Computing, 2023. Available at: https://www.R-project.org/ (accessed 06.10.2026).

34. Virtanen P., Gommers R., Oliphant T.E., et al. SciPy 1.0: fundamental algorithms for scientific computing in Python. Nature Methods. 2020;17(3):261–272. DOI: 10.1038/s41592-019-0686-2

35. GOST 25100-2020. Grunty. Klassifikatsiya [Soils. Classification]. Moscow: Standartinform, 2020. (In Russ.)

36. Seed H.B., Woodward R.J. Jr., Lundgren R. Prediction of swelling potential for compacted clays. Journal of the Soil Mechanics and Foundations Division. 1962;88(3):53–87. DOI: 10.1061/JSFEAQ.0000431

37. Sposito G. The Chemistry of Soils. 2nd ed. New York: Oxford University Press, 2008. 329 p.

38. Kelleners T.J., Robinson D.A., Shouse P.J., et al. Frequency dependence of the complex permittivity and its impact on dielectric sensor calibration in soils. Soil Science Society of America Journal. 2005;69(1):67–76. DOI: 10.2136/sssaj2005.0067a

39. GOST 12248.6-2020. Grunty. Metod opredeleniya nabukhaniya i usadki [Soils. Method for determination of swelling and shrinkage]. Moscow: Standartinform, 2020. (In Russ.)

СПИСОК ЛИТЕРАТУРЫ

1. Hillel, D. Introduction to Environmental Soil Physics / D. Hillel. Amsterdam : Elsevier Academic Press, 2004. 494 p.

2. Brady, N. C. Elements of the Nature and Properties of Soils / N. C. Brady, R. R. Weil. 3rd ed. Upper Saddle River, NJ : Pearson Prentice Hall, 2010. 614 p.

3. CS616 and CS625 Water Content Reflectometers : Product Manual. Rev. 01/2026. Logan, UT : Campbell Scientific, Inc., 2026. URL: https://s.campbellsci.com/documents/us/manuals/cs616.pdf (дата обращения: 06.10.2026).

4. Topp, G. C. Electromagnetic determination of soil water content : measurements in coaxial transmission lines / G. C. Topp, J. L. Davis, A. P. Annan // Water Resources Research. 1980. Vol. 16, No. 3. P. 574–582. DOI: 10.1029/WR016i003p00574.

5. A review of advances in dielectric and electrical conductivity measurement in soils using time domain reflectometry / D. A. Robinson, S. B. Jones, J. M. Wraith [et al.] // Vadose Zone Journal. 2003. Vol. 2, No. 4. P. 444–475. DOI: 10.2136/vzj2003.4440.

6. Bosch, D. D. Comparison of capacitance-based soil water probes in coastal plain soils / D. D. Bosch // Vadose Zone Journal. 2004. Vol. 3, No. 4. P. 1380–1389. DOI: 10.2136/vzj2004.1380.

7. Varble, J. L. Performance evaluation and calibration of soil water content and potential sensors for agricultural soils in eastern Colorado / J. L. Varble, J. L. Chávez // Agricultural Water Management. 2011. Vol. 101, No. 1. P. 93–106. DOI: 10.1016/j.agwat.2011.09.007.

8. Evaluation of standard calibration functions for eight electromagnetic soil moisture sensors / C. M. P. Vaz, S. B. Jones, M. Meding, M. Tuller // Vadose Zone Journal. 2013. Vol. 12, No. 2. P. 1–16. DOI: 10.2136/vzj2012.0160.

9. Почвы Краснодарского края, их использование и охрана / В. Ф. Вальков, Ю. А. Штомпель, И. Т. Трубилин [и др.]. Ростов-на-Дону : Изд-во СКНЦ ВШ, 1996. 192 с.

10. Качинский, Н. А. Механический и микроагрегатный состав почвы, методы его изучения / Н. А. Качинский. Москва : Изд-во АН СССР, 1958. 192 с.

11. Шеин, Е. В. Курс физики почв / Е. В. Шеин. Москва : Изд-во МГУ, 2005. 430 с.

12. Calibration of time domain reflectometry for water content measurement using a composite dielectric approach / K. Roth, R. Schulin, H. Flühler, W. Attinger // Water Resources Research. 1990. Vol. 26, No. 10. P. 2267–2273. DOI: 10.1029/WR026i010p02267.

13. Nelson, J. D. Expansive Soils : Problems and Practice in Foundation and Pavement Engineering / J. D. Nelson, D. J. Miller. New York : Wiley, 1992. 259 p.

14. Impact of soil profile mineralogy on the elemental composition of Chardonnay grapes and wines in the Anapa region / Z. Temerdashev, A. Abakumov, M. Bolshov [et al.] // Beverages. 2026. Vol. 12, No. 1. Art. 1. DOI: 10.3390/beverages12010001.

15. Адлер, Ю. П. Планирование эксперимента при поиске оптимальных условий / Ю. П. Адлер, Е. В. Маркова, Ю. В. Грановский. 2-е изд., перераб. и доп. Москва : Наука, 1976. 279 с.

16. Diagnosis and Improvement of Saline and Alkali Soils / ed. L. A. Richards. Washington, DC : USDA, 1954. 160 p. (Agriculture Handbook ; No. 60).

17. Классификация и диагностика почв СССР / В. В. Егоров, В. М. Фридланд, Е. Н. Иванова [и др.]. Москва : Колос, 1977. 223 с.

18. ГОСТ 17.4.4.02–2017. Охрана природы. Почвы. Методы отбора и подготовки проб для химического, бактериологического, гельминтологического анализа. Москва : Стандартинформ, 2018.

19. ГОСТ 17.4.3.01–2017. Охрана природы. Почвы. Общие требования к отбору проб. Москва : Стандартинформ, 2018.

20. ГОСТ 5180–2015. Грунты. Методы лабораторного определения физических характеристик. Москва : Стандартинформ, 2016.

21. ISO/IEC 17025:2017. General requirements for the competence of testing and calibration laboratories. Geneva : ISO, 2017. 30 p.

22. Mitchell, J. K. Fundamentals of Soil Behavior / J. K. Mitchell, K. Soga. 3rd ed. Hoboken, NJ : John Wiley & Sons, 2005. 577 p.

23. ГОСТ Р 51568–99 (ИСО 3310-1–90). Сита лабораторные из металлической проволочной сетки. Технические условия. Москва : Изд-во стандартов, 2000.

24. Ганжара, Н. Ф. Практикум по почвоведению / Н. Ф. Ганжара, Б. А. Борисов, Р. Ф. Байбеков ; под ред. Н. Ф. Ганжары. Москва : Агроконсалт, 2002. 280 с.

25. ГОСТ 12536–2014. Грунты. Методы лабораторного определения гранулометрического (зернового) и микроагрегатного состава. Москва : Стандартинформ, 2015.

26. ГОСТ 28268–89. Почвы. Методы определения влажности, максимальной гигроскопической влажности и влажности устойчивого завядания растений. Москва : Изд-во стандартов, 1989.

27. ГОСТ 26423–85. Почвы. Методы определения удельной электрической проводимости, pH и плотного остатка водной вытяжки. Москва : Изд-во стандартов, 1985.

28. Slavich, P. G. Estimating the electrical conductivity of saturated paste extracts from 1:5 soil:water suspensions and texture / P. G. Slavich, G. H. Petterson // Australian Journal of Soil Research. 1993. Vol. 31, No. 1. P. 73–81. DOI: 10.1071/SR9930073.

29. Hignett, C. Direct and surrogate measures of soil water content / C. Hignett, S. Evett // Field Estimation of Soil Water Content : A Practical Guide to Methods, Instrumentation and Sensor Technology. Vienna : IAEA, 2008. P. 1–22. (Training Course Series ; No. 30).

30. Datta, S. Soil water sensors for irrigation scheduling in the United States : a systematic review of literature / S. Datta, S. Taghvaeian // Agricultural Water Management. 2023. Vol. 278. Art. 108148. DOI: 10.1016/j.agwat.2023.108148.

31. Воронин, А. Д. Основы физики почв / А. Д. Воронин. Москва : Изд-во МГУ, 1986. 243, [1] с.

32. Качинский, Н. А. Физика почвы. Ч. 1 / Н. А. Качинский. Москва : Высшая школа, 1965. 323 с.

33. R: A language and environment for statistical computing / R Core Team. Vienna : R Foundation for Statistical Computing, 2023. URL: https://www.R-project.org/ (дата обращения: 06.10.2026).

34. SciPy 1.0: fundamental algorithms for scientific computing in Python / P. Virtanen, R. Gommers, T. E. Oliphant [et al.] // Nature Methods. 2020. Vol. 17, No. 3. P. 261–272. DOI: 10.1038/s41592-019-0686-2.

35. ГОСТ 25100–2020. Грунты. Классификация. Москва : Стандартинформ, 2020.

36. Seed, H. B. Prediction of swelling potential for compacted clays / H. B. Seed, R. J. Woodward Jr., R. Lundgren // Journal of the Soil Mechanics and Foundations Division. 1962. Vol. 88, No. 3. P. 53–87. DOI: 10.1061/JSFEAQ.0000431.

37. Sposito, G. The Chemistry of Soils / G. Sposito. 2nd ed. New York : Oxford University Press, 2008. 329 p.

38. Frequency dependence of the complex permittivity and its impact on dielectric sensor calibration in soils / T. J. Kelleners, D. A. Robinson, P. J. Shouse [et al.] // Soil Science Society of America Journal. 2005. Vol. 69, No. 1. P. 67–76. DOI: 10.2136/sssaj2005.0067a.

39. ГОСТ 12248.6–2020. Грунты. Метод определения набухания и усадки. Москва : Стандартинформ, 2020.

СВЕДЕНИЯ ОБ АВТОРАХ

Вербицкий Артём Юрьевич, аспирант кафедры строительства и эксплуатации водохозяйственных объектов, Кубанский государственный аграрный университет имени И. Т. Трубилина, ул. Калинина, 13, г. Краснодар, 350044, Российская Федерация, artem.verbitsk@yandex.ru

Приходько Игорь Александрович, канд. техн. наук, доцент, заведующий кафедрой строительства и эксплуатации водохозяйственных объектов, Кубанский государственный аграрный университет имени И. Т. Трубилина, ул. Калинина, 13, г. Краснодар, 350044, Российская Федерация, prihodkoigor2012@yandex.ru

Болдырева Леся Михайловна, канд. техн. наук, доцент кафедры строительства и эксплуатации водохозяйственных объектов, Кубанский государственный аграрный университет имени И. Т. Трубилина, ул. Калинина, 13, г. Краснодар, 350044, Российская Федерация, boldyreva.l@edu.kubsau.ru

Кесафоти Христофор Евстафьевич, ассистент кафедры строительства и эксплуатации водохозяйственных объектов, Кубанский государственный аграрный университет имени И. Т. Трубилина, ул. Калинина, 13, г. Краснодар, 350044, Российская Федерация, mr.kesafoti@mail.ru

INFORMATION ABOUT THE AUTHORS

Artem Yu. Verbitsky, Postgraduate Student, Department of Construction and Operation of Water Management Facilities, Kuban State Agrarian University named after I. T. Trubilin, 13 Kalinina St., Krasnodar, 350044, Russian Federation, artem.verbitsk@yandex.ru; SPIN-code: 8489-3868; ORCID: https://orcid.org/0000-0001-6930-2662; RSCI AuthorID: 1195740

Igor A. Prikhodko, Cand. Sci. (Eng.), Associate Professor, Head of the Department of Construction and Operation of Water Management Facilities, Kuban State Agrarian University named after I. T. Trubilin, 13 Kalinina St., Krasnodar, 350044, Russian Federation, prihodkoigor2012@yandex.ru; SPIN-code: 4011-7185; ORCID: https://orcid.org/0000-0003-4855-0434; RSCI AuthorID: 970218; WoS ResearcherID: AAH-1647-2021; Scopus Author ID: 57214098822

Lesya M. Boldyreva, Cand. Sci. (Eng.), Associate Professor, Department of Construction and Operation of Water Management Facilities, Kuban State Agrarian University named after I. T. Trubilin, 13 Kalinina St., Krasnodar, 350044, Russian Federation, boldyreva.l@edu.kubsau.ru; SPIN-code: 3928-3196; ORCID: https://orcid.org/0009-0002-0869-9859; Scopus Author ID: 60214676600

Khristofor E. Kesafoti, Assistant Lecturer, Department of Construction and Operation of Water Management Facilities, Kuban State Agrarian University named after I. T. Trubilin, 13 Kalinina St., Krasnodar, 350044, Russian Federation, mr.kesafoti@mail.ru; SPIN-code: 1186-0294; ORCID: https://orcid.org/0009-0008-7468-865X